## calibrations oct-2024
Serial / material | Blank | AVERAGE RAW READING 400ppb RWT DYE(40ug/L chla equiv) | Calibration curve | Gain setting | Calibration liquid | date
2103415 Titanium | 0 | 838.9091 | 0.047681 | 10x | 400ppb rhodamine WT dye | 2024-10-01
2180450 plastic | 0 | 2020.8 | 0.019794 | 10x | 400ppb rhodamine WT dye | 2024-10-01
2180556 plastic | 0 | 2662.5 | 0.015023 | 10x | 400ppb rhodamine WT dye | 2024-10-01

## Temperature cal apr-2025
Temperature | AP# | Test# | Raw Value |  | 1 | s1 | s2 | s3 | Temperature
30 | 1 | 1 | 2084.7 |  |  | 2084.8333 | 2115.4 | 2062.8667 | 30
| 1 | 2 | 2084.6 |  |  | 1989.3333 | 2015 | 1966.4 | 26
| 1 | 3 | 2085.2 |  |  | 1807.4333 | 1824.7667 | 1783.9667 | 19.25
| 2 | 1 | 2115.9 | # |  | 1678.5 | 1690.1333 | 1653.3 | 15
| 2 | 2 | 2114.9 | #
| 2 | 3 | 2115.4 | #
| 3 | 1 | 2063.3
| 3 | 2 | 2062.8
| 3 | 3 | 2062.5
41.025 | 1 | 1 | 2300.2
41.025 | 1 | 2 | 2298.7
41.025 | 1 | 3 | 2299.6
41.025 | 2 | 1 | 2334.6
41.025 | 2 | 2 | 2336.5
41.025 | 2 | 3 | 2337.2
41.025 | 3 | 1 | 2280.7
41.025 | 3 | 2 | 2278.8
41.025 | 3 | 3 | 2279.1
26 | 1 | 1 | 1989.1
26 | 1 | 2 | 1989.4
26 | 1 | 3 | 1989.5
26 | 2 | 1 | 2015
26 | 2 | 2 | 2014.2
26 | 2 | 3 | 2015.8
26 | 3 | 1 | 1966.4
26 | 3 | 2 | 1966.4
26 | 3 | 3 | 1966.4
19.25 | 1 | 1 | 1806.7
| 1 | 2 | 1808.2
| 1 | 3 | 1807.4
19.25 | 2 | 1 | 1825
| 2 | 2 | 1824.7
| 2 | 3 | 1824.6
19.25 | 3 | 1 | 1784.7
| 3 | 2 | 1783.5
| 3 | 3 | 1783.7 | [7 empty] | 3
15 | 1 | 1 | 1677.8
| 1 | 2 | 1678.6
| 1 | 3 | 1679.1
15 | 2 | 1 | 1689.2
| 2 | 2 | 1690.2
| 2 | 3 | 1691 | [6 empty] | Te
15 | 3 | 1 | 1653.4
| 3 | 2 | 1652.9
| 3 | 3 | 1653.6

## 2180450
Time | Probe_TempRaw | Probe_TempCal | Condraw | CondCal | SpCond | Salinity | TurbRaw | TurbCal | TurbManu | ChlRaw | ChlVolts | ChlCal | CDOMRaw | CDOMVolts | CDOMCal | CDOMChlEQ | ChlAdj | TempRaw | TempCal
2024-10-17 17:37 | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | 1984 | 1.5984 | 10.9112 | 0 | 0 | 0 | 0 | 10.9112 | 800 | 6.86
2024-10-17 17:37 | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | 1972 | 1.5888 | 10.8428 | 0 | 0 | 0 | 0 | 10.8428 | 792 | 6.7752
2024-10-17 17:38 | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | 2064 | 1.6629 | 11.3672 | 0 | 0 | 0 | 0 | 11.3672 | 793 | 6.7858
2024-10-17 17:38 | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | 2050 | 1.6516 | 11.2874 | 0 | 0 | 0 | 0 | 11.2874 | 781 | 6.6586
2024-10-17 17:39 | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | 2010 | 1.6194 | 11.0594 | 0 | 0 | 0 | 0 | 11.0594 | 790 | 6.754
2024-10-17 17:39 | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | 2035 | 1.6395 | 11.2019 | 0 | 0 | 0 | 0 | 11.2019 | 797 | 6.8282
2024-10-17 17:40 | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | 1998 | 1.6097 | 10.991 | 0 | 0 | 0 | 0 | 10.991 | 791 | 6.7646
2024-10-17 17:40 | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | 2018 | 1.6258 | 11.105 | 0 | 0 | 0 | 0 | 11.105 | 783 | 6.6798
2024-10-17 17:41 | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | 2079 | 1.675 | 11.4527 | 0 | 0 | 0 | 0 | 11.4527 | 809 | 6.9554
2024-10-17 17:41 | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | 1998 | 1.6097 | 10.991 | 0 | 0 | 0 | 0 | 10.991 | 790 | 6.754

## 2180556
Time | Probe_TempRaw | Probe_TempCal | Condraw | CondCal | SpCond | Salinity | TurbRaw | TurbCal | TurbManu | ChlRaw | ChlVolts | ChlCal | CDOMRaw | CDOMVolts | CDOMCal | CDOMChlEQ | ChlAdj | TempRaw | TempCal
2024-10-17 17:58 | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | 2627 | 2.1165 | 14.5763 | 0 | 0 | 0 | 0 | 14.5763 | 824 | 7.1144
2024-10-17 17:58 | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | 2603 | 2.0971 | 14.4395 | 0 | 0 | 0 | 0 | 14.4395 | 820 | 7.072
2024-10-17 17:59 | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | 2614 | 2.106 | 14.5022 | 0 | 0 | 0 | 0 | 14.5022 | 819 | 7.0614
2024-10-17 17:59 | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | 2695 | 2.1713 | 14.9639 | 0 | 0 | 0 | 0 | 14.9639 | 816 | 7.0296
2024-10-17 18:00 | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | 2639 | 2.1261 | 14.6447 | 0 | 0 | 0 | 0 | 14.6447 | 825 | 7.125
2024-10-17 18:02 | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | 2651 | 2.1358 | 14.7131 | 0 | 0 | 0 | 0 | 14.7131 | 802 | 6.8812
2024-10-17 18:03 | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | 2683 | 2.1616 | 14.8955 | 0 | 0 | 0 | 0 | 14.8955 | 801 | 6.8706
2024-10-17 18:04 | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | 2721 | 2.1922 | 15.1121 | 0 | 0 | 0 | 0 | 15.1121 | 797 | 6.8282
2024-10-17 18:05 | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | 2706 | 2.1801 | 15.0266 | 0 | 0 | 0 | 0 | 15.0266 | 789 | 6.7434
2024-10-17 18:06 | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | Fail | 2686 | 2.164 | 14.9126 | 0 | 0 | 0 | 0 | 14.9126 | 774 | 6.5844

## TItanium raw
Time | Probe_TempRaw | Probe_TempCal | Condraw | CondCal | SpCond | Salinity | TurbRaw | TurbCal | TurbManu | ChlRaw | ChlVolts | ChlCal | CDOMRaw | CDOMVolts | CDOMCal | CDOMChlEQ | ChlAdj | TempRaw | TempCal
2024-10-17 19:18 | Fail | Fail | Fail | Fail | Fail | Fail | 1534 | 1.2359 | 8.3462 | 883 | 0.7114013671874999 | 28.4561 | 0.9713 | 7.374899999999999 | 1409 | 13.3154
2024-10-17 19:19 | Fail | Fail | Fail | Fail | Fail | Fail | 1517 | 1.2221923828124999 | 8.2493 | 879 | 0.70818 | 28.3271 | 0.9669000000000001 | 7.2824 | 818 | 7.0508
2024-10-17 19:20 | Fail | Fail | Fail | Fail | Fail | Fail | 1528 | 1.2311 | 8.312 | 851 | 0.6856201171874999 | 27.4248 | 0.9361 | 7.3759 | 1459 | 13.845400000000001
2024-10-17 19:20 | Fail | Fail | Fail | Fail | Fail | Fail | 1517 | 1.2221923828124999 | 8.2493 | 893 | 0.7194580078124999 | 28.778318080000002 | 0.9823000000000001 | 7.2669999999999995 | 822 | 7.0932
2024-10-17 19:20 | Fail | Fail | Fail | Fail | Fail | Fail | 1518 | 1.223 | 8.254999999999999 | 846 | 0.6815917968749999 | 27.263669760000003 | 0.9306000000000001 | 7.324399999999999 | 1470 | 13.962
2024-10-17 19:21 | Fail | Fail | Fail | Fail | Fail | Fail | 1536 | 1.2374999999999998 | 8.3576 | 826 | 0.6654785156249999 | 26.6191 | 0.9086000000000001 | 7.449 | 824 | 7.114400000000001
2024-10-17 19:22 | Fail | Fail | Fail | Fail | Fail | Fail | 1517 | 1.2221923828124999 | 8.2493 | 809 | 0.65178 | 26.0713 | 0.8899 | 7.3594 | 1459 | 13.845400000000001
2024-10-17 19:24 | Fail | Fail | Fail | Fail | Fail | Fail | 1491 | 1.2012 | 8.101099999999999 | 804 | 0.64775 | 25.9102 | 0.8844000000000001 | 7.216699999999999 | 1092 | 9.955200000000001
2024-10-17 19:25 | Fail | Fail | Fail | Fail | Fail | Fail | 1542 | 1.2423339843749999 | 8.3918 | 834 | 0.67192 | 26.877 | 0.9174000000000001 | 7.4744 | 857 | 7.464199999999999
2024-10-17 19:34 | Fail | Fail | Fail | Fail | Fail | Fail | 1533 | 1.2351 | 8.3405 | 793 | 0.6388916015624999 | 25.5557 | 0.8723000000000001 | 7.4682 | 1302 | 12.1812
2024-10-17 19:35 | Fail | Fail | Fail | Fail | Fail | Fail | 1561 | 1.2576416015624998 | 8.5001 | 810 | 0.65259 | 26.1035 | 0.891 | 7.6091 | 1048 | 9.488800000000001
2024-10-17 19:36 | Fail | Fail | Fail | Fail | Fail | Fail | 4095 | 3.2992 | 22.9439 | 61 | 0.049145507812499996 | 1.9658 | 0.0671 | 22.8768 | 1223 | 11.343800000000002
2024-10-17 19:36 | Fail | Fail | Fail | Fail | Fail | Fail | 4095 | 3.2992 | 22.9439 | 82 | 0.066064 | 2.6426 | 0.0902 | 22.8537 | 1201 | 11.110600000000002
838.9091

## Calibrations dec-2024
Sensor Serial number | aqua pi # | Gradient | Intercept | gain | Equation | Test
| AP01 | 0.026735 | -2.1026 | 1 |  | 1500 | 38.0005
|  | 0.026735 | -2.1026 | 10 | y = 0.0267354105x - 2.1025729662 R² = 0.9977543048 | 1500 | 38.0005
|  | 0.02307 | -0.28051 | 100 | y = 0.0230696519x - 0.2805106216 R² = 0.9998647614 | 1500 | 34.324
| AP02 | 0.020155 | 0.0506 | 1 | N/A use 10x | 1500 | 30.2838
|  | 0.020155 | 0.0506 | 10 | y = 0.0201554957x + 0.0506003903 R² = 0.9987098006 | 1500 | 30.2838
|  | 0.018837 | -0.10112 | 100 | y = 0.0188365722x - 0.1011194574 R² = 0.9992419959 | 1500 | 28.1537
| AP03 | 0.032164 | -0.42538 | 1 | N/A use 10x | 1280 | 40.7447
|  | 0.032164 | -0.42538 | 10 | y = 0.0321641372x - 0.4253764149 R² = 0.9995672940 | 1280 | 40.7447
|  | 0.029298 | -0.74183 | 100 | y = 0.0292977027x - 0.7418318107 R² = 0.9996098059 | 1280 | 36.7592

## Linear_AP03
Concentration | ChlRaw | ChlRawRangeCI | ChlVolts | ChlVoltsRangeCI | ChlCal | ChlCalRangeCI | Gain | ChlRaw 10x equivalent | Chl equivalent
800 | 2483.8929 | 4.8367 | 3.0328 | 0.0059056 | 37.3167 | 0.072663 | 10x | 2483.8929 | 80
800 | 2489.2447 | 4.5356 | 3.0394 | 0.005538 | 37.3971 | 0.068141 | 10x | 2489.2447 | 80
800 | 2487.4858 | 4.4064 | 3.0372 | 0.0053802 | 37.3707 | 0.066199 | 10x | 2487.4858 | 80
800 | 2487.5536 | 4.1311 | 3.0373 | 0.005044 | 37.3717 | 0.062063 | 10x | 2487.5536 | 80
800 | 2493.3203 | 4.085 | 3.0443 | 0.0049878 | 37.4583 | 0.061371 | 10x | 2493.3203 | 80
400 | 1284.4397 | 4.2521 | 1.5683 | 0.0051918 | 19.2967 | 0.063881 | 10x | 1284.4397 | 40
400 | 1280.344 | 4.6242 | 1.5633 | 0.0056461 | 19.2352 | 0.069471 | 10x | 1280.344 | 40
400 | 1277.9211 | 4.512 | 1.5603 | 0.0055092 | 19.1988 | 0.067786 | 10x | 1277.9211 | 40
400 | 1282.3143 | 4.7448 | 1.5657 | 0.0057935 | 19.2648 | 0.071284 | 10x | 1282.3143 | 40
400 | 1283.735 | 4.7696 | 1.5674 | 0.0058237 | 19.2862 | 0.071656 | 10x | 1283.735 | 40
200 | 639.4291 | 4.3911 | 0.78074 | 0.0053615 | 9.6064 | 0.065969 | 10x | 639.4291 | 20
200 | 642.117 | 4.2221 | 0.78403 | 0.0051552 | 9.6468 | 0.06343 | 10x | 642.117 | 20
200 | 641.9466 | 4.3283 | 0.78382 | 0.0052849 | 9.6443 | 0.065026 | 10x | 641.9466 | 20
200 | 642.6762 | 4.3379 | 0.78471 | 0.0052966 | 9.6552 | 0.06517 | 10x | 642.6762 | 20
200 | 641.9815 | 4.5065 | 0.78386 | 0.0055025 | 9.6448 | 0.067704 | 10x | 641.9815 | 20
100 | 303.6631 | 4.4908 | 0.37077 | 0.0054833 | 4.5621 | 0.067468 | 10x | 303.6631 | 10
100 | 305.7518 | 4.3178 | 0.37332 | 0.0052721 | 4.5935 | 0.064869 | 10x | 305.7518 | 10
100 | 304.3381 | 4.2673 | 0.3716 | 0.0052104 | 4.5722 | 0.06411 | 10x | 304.3381 | 10
100 | 304.3121 | 4.3426 | 0.37157 | 0.0053023 | 4.5718 | 0.06524 | 10x | 304.3121 | 10
100 | 305.9326 | 4.2937 | 0.37354 | 0.0052426 | 4.5962 | 0.064506 | 10x | 305.9326 | 10
50 | 1950.5248 | 4.3929 | 2.3816 | 0.0053637 | 2.9304 | 0.0065996 | 100x | 195.0525 | 5
50 | 1963.4234 | 4.3699 | 2.3973 | 0.0053356 | 2.9497 | 0.0065651 | 100x | 196.3423 | 5
50 | 1961.0496 | 4.3476 | 2.3944 | 0.0053084 | 2.9462 | 0.0065316 | 100x | 196.105 | 5
50 | 1958.1773 | 5.1587 | 2.3909 | 0.0062988 | 2.9419 | 0.0077502 | 100x | 195.8177 | 5
50 | 1952.3985 | 4.479 | 2.3839 | 0.0054688 | 2.9332 | 0.006729 | 100x | 195.2398 | 5
25 | 1110.8754 | 4.2338 | 1.3564 | 0.0051695 | 1.6689 | 0.0063606 | 100x | 111.0875 | 2.5
25 | 1113.5743 | 4.5058 | 1.3597 | 0.0055016 | 1.673 | 0.0067693 | 100x | 111.3574 | 2.5
25 | 1110.0213 | 4.7916 | 1.3553 | 0.0058506 | 1.6676 | 0.0071987 | 100x | 111.0021 | 2.5
25 | 1108.4304 | 4.9097 | 1.3534 | 0.0059948 | 1.6652 | 0.0073761 | 100x | 110.843 | 2.5
25 | 1108.1815 | 4.4393 | 1.3531 | 0.0054204 | 1.6649 | 0.0066694 | 100x | 110.8181 | 2.5
12.5 | 685.1135 | 4.4485 | 0.83652 | 0.0054316 | 1.0293 | 0.0066832 | 100x | 68.5113 | 1.25
12.5 | 676.349 | 4.3361 | 0.82582 | 0.0052944 | 1.0161 | 0.0065143 | 100x | 67.6349 | 1.25
12.5 | 680.695 | 4.6669 | 0.83113 | 0.0056982 | 1.0226 | 0.0070112 | 100x | 68.0695 | 1.25
12.5 | 682.0473 | 4.664 | 0.83278 | 0.0056947 | 1.0247 | 0.0070069 | 100x | 68.2047 | 1.25
12.5 | 680.7837 | 4.6017 | 0.83124 | 0.0056186 | 1.0228 | 0.0069133 | 100x | 68.0784 | 1.25
6.25 | 464.6833 | 4.341 | 0.56738 | 0.0053004 | 0.69812 | 0.0065217 | 100x | 46.4683 | 0.625
6.25 | 475.2324 | 4.8569 | 0.58026 | 0.0059302 | 0.71396 | 0.0072967 | 100x | 47.5232 | 0.625
6.25 | 467.4539 | 4.8149 | 0.57076 | 0.005879 | 0.70228 | 0.0072336 | 100x | 46.7454 | 0.625
6.25 | 472.9269 | 4.7638 | 0.57744 | 0.0058166 | 0.7105 | 0.0071568 | 100x | 47.2927 | 0.625
6.25 | 468.2482 | 4.2964 | 0.57173 | 0.0052459 | 0.70347 | 0.0064546 | 100x | 46.8248 | 0.625
3.125 | 352.581 | 4.5354 | 0.4305 | 0.0055377 | 0.5297 | 0.0068137 | 100x | 35.2581 | 0.3125
3.125 | 363.5177 | 4.3505 | 0.44386 | 0.005312 | 0.54613 | 0.006536 | 100x | 36.3518 | 0.3125
3.125 | 356.5316 | 4.7737 | 0.43533 | 0.0058286 | 0.53563 | 0.0071717 | 100x | 35.6532 | 0.3125
3.125 | 359.9217 | 4.2342 | 0.43946 | 0.00517 | 0.54073 | 0.0063613 | 100x | 35.9922 | 0.3125
3.125 | 355.3591 | 4.4349 | 0.43389 | 0.005415 | 0.53387 | 0.0066627 | 100x | 35.5359 | 0.3125
1.33 | 290.2632 | 4.3592 | 0.35441 | 0.0053226 | 0.43608 | 0.0065491 | 100x | 29.0263 | 0.133
1.33 | 284.2801 | 4.3576 | 0.34711 | 0.0053206 | 0.42709 | 0.0065466 | 100x | 28.428 | 0.133
1.33 | 286.2308 | 4.9151 | 0.34949 | 0.0060013 | 0.43002 | 0.0073841 | 100x | 28.6231 | 0.133
1.33 | 281.0638 | 4.1945 | 0.34318 | 0.0051214 | 0.42226 | 0.0063015 | 100x | 28.1064 | 0.133
1.33 | 282.2874 | 4.758 | 0.34467 | 0.0058096 | 0.42409 | 0.0071482 | 100x | 28.2287 | 0.133
0.267 | 290.2632 | 4.3592 | 0.35441 | 0.0053226 | 0.43608 | 0.0065491 | 100x | 29.0263 | 0.0267
0.267 | 284.2801 | 4.3576 | 0.34711 | 0.0053206 | 0.42709 | 0.0065466 | 100x | 28.428 | 0.0267
0.267 | 286.2308 | 4.9151 | 0.34949 | 0.0060013 | 0.43002 | 0.0073841 | 100x | 28.6231 | 0.0267
0.267 | 281.0638 | 4.1945 | 0.34318 | 0.0051214 | 0.42226 | 0.0063015 | 100x | 28.1064 | 0.0267
0.267 | 282.2874 | 4.758 | 0.34467 | 0.0058096 | 0.42409 | 0.0071482 | 100x | 28.2287 | 0.0267
0 | 237.6949 | 4.684 | 0.29023 | 0.0057191 | 0.3571 | 0.007037 | 100x | 23.7695 | 0
0 | 243.727 | 4.3733 | 0.29759 | 0.0053398 | 0.36616 | 0.0065702 | 100x | 24.3727 | 0
0 | 240.7313 | 4.6586 | 0.29393 | 0.0056881 | 0.36166 | 0.0069988 | 100x | 24.0731 | 0
0 | 244.0106 | 4.6289 | 0.29794 | 0.0056519 | 0.36659 | 0.0069542 | 100x | 24.4011 | 0
0 | 241.0887 | 4.2306 | 0.29437 | 0.0051655 | 0.3622 | 0.0063558 | 100x | 24.1089 | 0

## Linear_AP02
Concentration | ChlRaw | ChlRawRangeCI | ChlVolts | ChlVoltsRangeCI | ChlCal | ChlCalRangeCI | Gain | ChlRaw 10x equivalent | Chl equivalent
800 | 315.9835 | 3.6926 | 0.38582 | 0.0045086 | 47.4717 | 0.55475 | 1x | 3159.8354 | 80
800 | 318.4588 | 3.7863 | 0.38884 | 0.0046231 | 47.8436 | 0.56884 | 1x | 3184.5878 | 80
800 | 318.7536 | 3.5941 | 0.3892 | 0.0043884 | 47.8879 | 0.53996 | 1x | 3187.5357 | 80
800 | 329.9591 | 3.6424 | 0.40288 | 0.0044474 | 49.5713 | 0.54721 | 1x | 3299.5911 | 80
800 | 318.9283 | 3.8942 | 0.38941 | 0.0047548 | 47.9141 | 0.58504 | 1x | 3189.2832 | 80
400 | 1968.5286 | 3.3124 | 2.4036 | 0.0040444 | 29.5741 | 0.049764 | 10x | 1968.5286 | 40
400 | 1970.7714 | 3.7042 | 2.4063 | 0.0045228 | 29.6078 | 0.05565 | 10x | 1970.7714 | 40
400 | 1973.1929 | 3.7146 | 2.4093 | 0.0045356 | 29.6442 | 0.055807 | 10x | 1973.1929 | 40
400 | 1973.225 | 3.7922 | 2.4093 | 0.0046303 | 29.6447 | 0.056972 | 10x | 1973.225 | 40
400 | 1973.4036 | 3.5302 | 2.4095 | 0.0043104 | 29.6474 | 0.053036 | 10x | 1973.4036 | 40
200 | 1036.7071 | 8.4112 | 1.2658 | 0.01027 | 15.5749 | 0.12637 | 10x | 1036.7071 | 20
200 | 1031.9573 | 7.5146 | 1.26 | 0.0091753 | 15.5036 | 0.1129 | 10x | 1031.9573 | 20
200 | 1041.3333 | 11.3409 | 1.2715 | 0.013847 | 15.6444 | 0.17038 | 10x | 1041.3333 | 20
200 | 984.1071 | 4.0466 | 1.2016 | 0.0049409 | 14.7847 | 0.060794 | 10x | 984.1071 | 20
100 | 468.1286 | 3.9373 | 0.57159 | 0.0048075 | 7.0329 | 0.059152 | 10x | 468.1286 | 10
100 | 479.6201 | 4.1888 | 0.58562 | 0.0051145 | 7.2056 | 0.06293 | 10x | 479.6201 | 10
100 | 506.6214 | 5.8573 | 0.61859 | 0.0071518 | 7.6112 | 0.087997 | 10x | 506.6214 | 10
100 | 483.1935 | 4.0943 | 0.58998 | 0.0049991 | 7.2592 | 0.06151 | 10x | 483.1935 | 10
100 | 485.7849 | 5.0488 | 0.59314 | 0.0061646 | 7.2982 | 0.075851 | 10x | 485.7849 | 10
100 | 494.825 | 6.3454 | 0.60418 | 0.0077478 | 7.434 | 0.09533 | 10x | 494.825 | 10
100 | 493.1183 | 4.2654 | 0.6021 | 0.005208 | 7.4083 | 0.064081 | 10x | 493.1183 | 10
100 | 472.694 | 3.8014 | 0.57716 | 0.0046415 | 7.1015 | 0.05711 | 10x | 472.694 | 10
100 | 474.875 | 3.6044 | 0.57982 | 0.004401 | 7.1343 | 0.054151 | 10x | 474.875 | 10
50 | 2696.8495 | 3.8196 | 3.2929 | 0.0046637 | 4.0516 | 0.0057384 | 100x | 269.6849 | 5
50 | 2675.9239 | 6.02 | 3.2673 | 0.0073504 | 4.0202 | 0.0090441 | 100x | 267.5924 | 5
50 | 2706.0627 | 4.1135 | 3.3041 | 0.0050226 | 4.0654 | 0.0061799 | 100x | 270.6063 | 5
50 | 2711.871 | 4.1952 | 3.3112 | 0.0051224 | 4.0742 | 0.0063027 | 100x | 271.1871 | 5
50 | 2711.7554 | 4.1774 | 3.3111 | 0.0051006 | 4.074 | 0.0062759 | 100x | 271.1755 | 5
25 | 1402.7204 | 3.5758 | 1.7127 | 0.004366 | 2.1074 | 0.0053721 | 100x | 140.272 | 2.5
25 | 1400.9424 | 4.3735 | 1.7106 | 0.00534 | 2.1047 | 0.0065705 | 100x | 140.0942 | 2.5
25 | 1415.9018 | 4.1155 | 1.7288 | 0.0050251 | 2.1272 | 0.006183 | 100x | 141.5902 | 2.5
25 | 1352.3273 | 6.851 | 1.6512 | 0.0083651 | 2.0317 | 0.010293 | 100x | 135.2327 | 2.5
25 | 1400.4359 | 4.0528 | 1.7099 | 0.0049484 | 2.1039 | 0.0060887 | 100x | 140.0436 | 2.5
12.5 | 718.5393 | 5.7993 | 0.87734 | 0.007081 | 1.0795 | 0.0087126 | 100x | 71.8539 | 1.25
12.5 | 634.7 | 10.5439 | 0.77497 | 0.012874 | 0.95354 | 0.015841 | 100x | 63.47 | 1.25
12.5 | 728.5639 | 5.9732 | 0.88958 | 0.0072932 | 1.0946 | 0.0089737 | 100x | 72.8564 | 1.25
12.5 | 727.4982 | 5.0021 | 0.88828 | 0.0061076 | 1.093 | 0.007515 | 100x | 72.7498 | 1.25
12.5 | 735.498 | 6.2727 | 0.89804 | 0.0076589 | 1.105 | 0.0094237 | 100x | 73.5498 | 1.25
6.25 | 406.187 | 5.9558 | 0.49595 | 0.0072721 | 0.61023 | 0.0089477 | 100x | 40.6187 | 0.625
6.25 | 379.3929 | 6.3849 | 0.46324 | 0.0077959 | 0.56998 | 0.0095923 | 100x | 37.9393 | 0.625
6.25 | 382.7571 | 6.4819 | 0.46735 | 0.0079144 | 0.57503 | 0.0097381 | 100x | 38.2757 | 0.625
6.25 | 381.4038 | 7.1803 | 0.46569 | 0.0087671 | 0.573 | 0.010787 | 100x | 38.1404 | 0.625
6.25 | 415.0964 | 4.6079 | 0.50683 | 0.0056263 | 0.62362 | 0.0069227 | 100x | 41.5096 | 0.625
3.125 | 243.9964 | 5.3689 | 0.29792 | 0.0065554 | 0.36657 | 0.0080659 | 100x | 24.3996 | 0.3125
3.125 | 214.4363 | 6.2779 | 0.26183 | 0.0076653 | 0.32216 | 0.0094315 | 100x | 21.4436 | 0.3125
3.125 | 172.5305 | 9.8145 | 0.21066 | 0.011983 | 0.2592 | 0.014745 | 100x | 17.253 | 0.3125
3.125 | 208.2739 | 7.4359 | 0.2543 | 0.0090793 | 0.3129 | 0.011171 | 100x | 20.8274 | 0.3125
3.125 | 231.825 | 6.0925 | 0.28306 | 0.007439 | 0.34828 | 0.0091531 | 100x | 23.1825 | 0.3125
1.33 | 137.1137 | 4.679 | 0.16742 | 0.0057131 | 0.20599 | 0.0070295 | 100x | 13.7114 | 0.133
1.33 | 141.0676 | 4.7287 | 0.17224 | 0.0057738 | 0.21193 | 0.0071042 | 100x | 14.1068 | 0.133
1.33 | 173.3893 | 4.1619 | 0.21171 | 0.0050817 | 0.26049 | 0.0062527 | 100x | 17.3389 | 0.133
1.33 | 141.5205 | 4.9902 | 0.1728 | 0.006093 | 0.21261 | 0.007497 | 100x | 14.152 | 0.133
1.33 | 143.0143 | 5.7654 | 0.17462 | 0.0070395 | 0.21486 | 0.0086616 | 100x | 14.3014 | 0.133
0.267 | 78.3234 | 4.0628 | 0.095633 | 0.0049606 | 0.11767 | 0.0061037 | 100x | 7.8323 | 0.0267
0.267 | 70.7061 | 4.9232 | 0.086332 | 0.0060112 | 0.10623 | 0.0073963 | 100x | 7.0706 | 0.0267
0.267 | 70.6094 | 5.6617 | 0.086214 | 0.006913 | 0.10608 | 0.0085059 | 100x | 7.0609 | 0.0267
0.267 | 46.8107 | 7.0468 | 0.057156 | 0.0086041 | 0.070326 | 0.010587 | 100x | 4.6811 | 0.0267
0.267 | 51.7228 | 7.5103 | 0.063154 | 0.00917 | 0.077706 | 0.011283 | 100x | 5.1723 | 0.0267
0 | 36.5878 | 7.4928 | 0.044674 | 0.0091487 | 0.054968 | 0.011257 | 100x | 3.6588 | 0
0 | 24.2236 | 7.2104 | 0.029577 | 0.0088039 | 0.036392 | 0.010832 | 100x | 2.4224 | 0
0 | 31.2536 | 7.1016 | 0.038161 | 0.008671 | 0.046954 | 0.010669 | 100x | 3.1254 | 0
0 | 39.5164 | 7.4805 | 0.04825 | 0.0091337 | 0.059367 | 0.011238 | 100x | 3.9516 | 0
0 | 44.4659 | 7.2965 | 0.054293 | 0.0089091 | 0.066803 | 0.010962 | 100x | 4.4466 | 0
`

## Linear_AP01
Concentration | ChlRaw | ChlRawRangeCI | ChlVolts | ChlVoltsRangeCI | ChlCal | ChlCalRangeCI | Gain | ChlRaw 10x equivalent | Chl equivalent
800 | 2984.8249 | 11.6077 | 3.6445 | 0.014173 | 44.8424 | 0.17439 | 10x | 2984.8249 | 80
800 | 3045.7766 | 3.6853 | 3.7189 | 0.0044997 | 45.7581 | 0.055366 | 10x | 3045.7766 | 80
800 | 3050.1281 | 3.232 | 3.7242 | 0.0039463 | 45.8235 | 0.048556 | 10x | 3050.1281 | 80
800 | 3047.9617 | 3.8443 | 3.7216 | 0.0046939 | 45.791 | 0.057755 | 10x | 3047.9617 | 80
800 | 3049.4146 | 3.2814 | 3.7233 | 0.0040066 | 45.8128 | 0.049298 | 10x | 3049.4146 | 80
400 | 1642.5426 | 3.1211 | 2.0055 | 0.0038109 | 24.6767 | 0.04689 | 10x | 1642.5426 | 40
400 | 1650.9004 | 2.883 | 2.0158 | 0.0035202 | 24.8023 | 0.043313 | 10x | 1650.9004 | 40
400 | 1649.5426 | 3.2019 | 2.0141 | 0.0039096 | 24.7819 | 0.048104 | 10x | 1649.5426 | 40
400 | 1649.6277 | 3.3775 | 2.0142 | 0.0041239 | 24.7831 | 0.050741 | 10x | 1649.6277 | 40
400 | 1643.9786 | 3.3141 | 2.0073 | 0.0040465 | 24.6983 | 0.04979 | 10x | 1643.9786 | 40
200 | 835.9539 | 3.0709 | 1.0207 | 0.0037496 | 12.5589 | 0.046136 | 10x | 835.9539 | 20
200 | 837.9375 | 3.3845 | 1.0231 | 0.0041325 | 12.5887 | 0.050848 | 10x | 837.9375 | 20
200 | 826.4424 | 3.4656 | 1.0091 | 0.0042315 | 12.416 | 0.052065 | 10x | 826.4424 | 20
200 | 829.4319 | 3.253 | 1.0127 | 0.0039719 | 12.4609 | 0.048871 | 10x | 829.4319 | 20
200 | 830.8992 | 3.0472 | 1.0145 | 0.0037207 | 12.483 | 0.04578 | 10x | 830.8992 | 20
100 | 410.0854 | 3.1356 | 0.50071 | 0.0038286 | 6.1609 | 0.047108 | 10x | 410.0854 | 10
100 | 411.0498 | 3.2386 | 0.50189 | 0.0039543 | 6.1754 | 0.048655 | 10x | 411.0498 | 10
100 | 409.6348 | 3.0551 | 0.50016 | 0.0037303 | 6.1541 | 0.045898 | 10x | 409.6348 | 10
100 | 408.4007 | 3.0993 | 0.49866 | 0.0037843 | 6.1356 | 0.046563 | 10x | 408.4007 | 10
100 | 411.0248 | 3.1427 | 0.50186 | 0.0038372 | 6.175 | 0.047214 | 10x | 411.0248 | 10
50 | 2286.0993 | 3.3299 | 2.7913 | 0.0040658 | 3.4345 | 0.0050026 | 100x | 228.6099 | 5
50 | 2287.8916 | 3.5356 | 2.7935 | 0.004317 | 3.4372 | 0.0053118 | 100x | 228.7892 | 5
50 | 2283.6277 | 3.0937 | 2.7883 | 0.0037774 | 3.4308 | 0.0046478 | 100x | 228.3628 | 5
50 | 2288.4221 | 3.2196 | 2.7942 | 0.0039312 | 3.438 | 0.004837 | 100x | 228.8422 | 5
50 | 2285.4752 | 3.4442 | 2.7906 | 0.0042053 | 3.4336 | 0.0051744 | 100x | 228.5475 | 5
25 | 1204.7046 | 3.6211 | 1.4709 | 0.0044214 | 1.8099 | 0.0054401 | 100x | 120.4705 | 2.5
25 | 1203.8243 | 3.282 | 1.4699 | 0.0040073 | 1.8086 | 0.0049307 | 100x | 120.3824 | 2.5
25 | 1205.2776 | 3.1667 | 1.4716 | 0.0038665 | 1.8107 | 0.0047574 | 100x | 120.5278 | 2.5
25 | 1204.1303 | 3.3031 | 1.4702 | 0.0040331 | 1.809 | 0.0049624 | 100x | 120.413 | 2.5
25 | 1205.2801 | 3.5408 | 1.4716 | 0.0043233 | 1.8107 | 0.0053195 | 100x | 120.528 | 2.5
12.5 | 669.0641 | 4.1783 | 0.81693 | 0.0051017 | 1.0052 | 0.0062773 | 100x | 66.9064 | 1.25
12.5 | 671.9113 | 3.3146 | 0.8204 | 0.0040471 | 1.0094 | 0.0049796 | 100x | 67.1911 | 1.25
12.5 | 670.4872 | 3.5027 | 0.81867 | 0.0042767 | 1.0073 | 0.0052622 | 100x | 67.0487 | 1.25
12.5 | 666.9424 | 3.8536 | 0.81434 | 0.0047053 | 1.002 | 0.0057895 | 100x | 66.6942 | 1.25
12.5 | 665.961 | 3.2386 | 0.81314 | 0.0039544 | 1.0005 | 0.0048655 | 100x | 66.5961 | 1.25
6.25 | 403.1317 | 3.4072 | 0.49222 | 0.0041602 | 0.60564 | 0.0051188 | 100x | 40.3132 | 0.625
6.25 | 401.7917 | 3.7367 | 0.49059 | 0.0045625 | 0.60363 | 0.0056138 | 100x | 40.1792 | 0.625
6.25 | 392.0284 | 5.2153 | 0.47867 | 0.0063679 | 0.58896 | 0.0078352 | 100x | 39.2028 | 0.625
6.25 | 400.9847 | 4.2118 | 0.4896 | 0.0051426 | 0.60242 | 0.0063275 | 100x | 40.0985 | 0.625
6.25 | 405.5957 | 4.709 | 0.49523 | 0.0057497 | 0.60935 | 0.0070746 | 100x | 40.5596 | 0.625
3.125 | 269.41 | 4.1365 | 0.32895 | 0.0050507 | 0.40475 | 0.0062145 | 100x | 26.941 | 0.3125
3.125 | 263.5694 | 3.3602 | 0.32182 | 0.0041028 | 0.39597 | 0.0050482 | 100x | 26.3569 | 0.3125
3.125 | 267.4122 | 4.3265 | 0.32651 | 0.0052827 | 0.40175 | 0.0065 | 100x | 26.7412 | 0.3125
3.125 | 265.1702 | 3.3666 | 0.32377 | 0.0041106 | 0.39838 | 0.0050578 | 100x | 26.517 | 0.3125
3.125 | 269.6738 | 3.6566 | 0.32927 | 0.0044648 | 0.40514 | 0.0054935 | 100x | 26.9674 | 0.3125
1.33 | 170.1037 | 4.4033 | 0.2077 | 0.0053765 | 0.25555 | 0.0066154 | 100x | 17.0104 | 0.133
1.33 | 158.5125 | 3.3727 | 0.19354 | 0.0041181 | 0.23814 | 0.005067 | 100x | 15.8512 | 0.133
1.33 | 154.8936 | 3.8849 | 0.18913 | 0.0047435 | 0.2327 | 0.0058365 | 100x | 15.4894 | 0.133
1.33 | 179.6838 | 3.8239 | 0.21939 | 0.0046689 | 0.26995 | 0.0057448 | 100x | 17.9684 | 0.133
0.267 | 124.8794 | 3.3973 | 0.15248 | 0.0041481 | 0.18761 | 0.0051039 | 100x | 12.4879 | 0.0267
0.267 | 135.5294 | 3.3702 | 0.16548 | 0.004115 | 0.20361 | 0.0050632 | 100x | 13.5529 | 0.0267
0.267 | 125.3286 | 3.2548 | 0.15303 | 0.0039742 | 0.18829 | 0.0048899 | 100x | 12.5329 | 0.0267
0.267 | 132.3227 | 3.3271 | 0.16157 | 0.0040624 | 0.19879 | 0.0049985 | 100x | 13.2323 | 0.0267
0.267 | 131.1097 | 3.6596 | 0.16009 | 0.0044683 | 0.19697 | 0.005498 | 100x | 13.111 | 0.0267
0 | 116.7046 | 3.9016 | 0.1425 | 0.0047639 | 0.17533 | 0.0058616 | 100x | 11.6705 | 0
0 | 112.1779 | 3.7048 | 0.13697 | 0.0045235 | 0.16853 | 0.0055658 | 100x | 11.2178 | 0
0 | 115.6595 | 4.5299 | 0.14122 | 0.005531 | 0.17376 | 0.0068054 | 100x | 11.5659 | 0
0 | 118.163 | 3.3805 | 0.14428 | 0.0041276 | 0.17752 | 0.0050787 | 100x | 11.8163 | 0
0 | 113.8617 | 3.9413 | 0.13903 | 0.0048123 | 0.17106 | 0.0059212 | 100x | 11.3862 | 0
[8 empty] | 0 | 0
`

## june 2025
Original_DateTime | Interval_Timestamp | AP01_DEC | ap01_10xequiv | AP01GAIN | Test-cl-example-ap02.ChlRaw |  | Test-cl-example-ap02.ChlGain | Test-cl-example-ap03.ChlRaw |  | Test-cl-example-ap03.ChlGain | test conc. | equiv conc
06/01/2025, 17:02:52 | 06/01/2025, 17:00:00 | 835.9539 | 835.9539 | 10x | 1008.5376 | 1008.5376 | 10x | 822.8759 | 822.8759 | 10x | 20 | 200
06/01/2025, 17:03:31 | 06/01/2025, 17:00:00 | 826.4424 | 826.4424 | 10x | 1012.3623 | 1012.3623 | 10x | 822.8759 | 822.8759 | 10x | 20 | 200
06/01/2025, 17:04:09 | 06/01/2025, 17:00:00 | 830.8992 | 830.8992 | 10x | 1010.1185 | 1010.1185 | 10x | 822.8759 | 822.8759 | 10x | 20 | 200
06/01/2025, 17:06:22 | 06/01/2025, 17:00:00 | 2286.0993 | 228.6099 | 100x | 2731.9331 | 273.1933 | 100x | 2359.6868 | 235.9687 | 100x | 5 | 50
06/01/2025, 17:07:02 | 06/01/2025, 17:00:00 | 2287.8916 | 228.7892 | 100x | 2726.2401 | 272.624 | 100x | 2364.3689 | 236.4369 | 100x | 5 | 50 |  | c1 | v1 | c2 | v2 | Vwater
06/01/2025, 17:07:42 | 06/01/2025, 17:15:00 | 2283.6277 | 228.3628 | 100x | 2724.6 | 272.46 | 100x | 2363.7456 | 236.3746 | 100x | 5 | 50 |  | 800 | 62.5 | 200 | 250 | 187.5
CALIBRATION 0 DATA |  | 1401.3157 | 140.1316 | 100x | 1445.5149 | 144.5515 | 100x | 998.0754 | 99.8075 | 100x | 0 | 0 |  | 200 | 62.5 | 50 | 250 | 187.5
CALIBRATION 0 DATA |  | 1401.3157 | 140.1316 | 100x | 1445.5149 | 144.5515 | 100x | 998.0754 | 99.8075 | 100x | 0 | 0
CALIBRATION 0 DATA |  | 1401.3157 | 140.1316 | 100x | 1445.5149 | 144.5515 | 100x | 998.0754 | 99.8075 | 100x | 0 | 0
u